## SCROD Write Registers
Register Number | (same in hex) | Col | Row | Ch | Name | Comments | Bit Mappings
0 | 0000 | N/A | N/A | N/A | LEDS | LEDs. Display SCROD ID on startup. FET must be installed properly. LED switch must be active. | One bit per LED
1 | 0001 | N/A | N/A | N/A | I2C_BUS[0]_Write | I2C write for SCROD EEPROM and temperature sensor on SCROD | Bits 15:13 unused, Bit 12: Send stop, Bit 11: Send acknowledge, Bit 10: Read byte, Bit 9: Send byte, Bit 8: Send start, Bit 7:0 Byte to send
2 | 0002 | N/A | N/A | N/A | I2C_BUS[1]_Write | I2C write for SCROD fiber transceiver 0 | Bits 15:13 unused, Bit 12: Send stop, Bit 11: Send acknowledge, Bit 10: Read byte, Bit 9: Send byte, Bit 8: Send start, Bit 7:0 Byte to send
3 | 0003 | N/A | N/A | N/A | I2C_BUS[2]_Write | I2C write for SCROD fiber transceiver 1 | Bits 15:13 unused, Bit 12: Send stop, Bit 11: Send acknowledge, Bit 10: Read byte, Bit 9: Send byte, Bit 8: Send start, Bit 7:0 Byte to send
4 | 0004 | N/A | N/A | N/A | Trigger interfaces | Trigger polarity, set column/row for scaler monitoring | Bit 15: trigger polarity (1 for rising edge, 0 for falling), Bits 14:4: unused, Bits 3:2: col select, Bits 1:0: row select (col/row selects choose which trigger scalers to read)
5 | 0005 | G | G | G | TRGBIAS | Internal ASIC trigger comparator bias, shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
6 | 0006 | G | G | G | TRGBIAS2 | Internal ASIC trigger comparator bias for reference channel, shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
7 | 0007 | G | G | G | TRGTHREF | Internal ASIC trigger threshold voltage for reference channel, shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
8 | 0008 | G | G | G | ISEL | Internal ASIC wilkinson ramp current source bias, shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
9 | 0009 | G | G | G | SBBIAS | Internal ASIC bias for the buffers that distribute the Wilkinson ramp, shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
10 | 000A | G | G | G | PUBIAS | Internal ASIC bias for pullup in wilkinson comparators, shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
11 | 000B | G | G | G | CMPBIAS | Internal ASIC bias for wilkinson comparators, shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
12 | 000C | G | G | G | DAC_BUF_BIASES | Internal ASIC biases for VBDbias, WBDbias, TCBbias, THDbias, Tbbias, TRGDbias, PDDbias, PUDbias, SBDbias, VDDbias, shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
13 | 000D | 0 | 0 | 0 | TRG_THRESH | Internal ASIC channel trigger thresholds - unique channel-to-channel | Bits 15:12 unused, Bits 11:0 DAC value
14 | 000E | 0 | 0 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
15 | 000F | 0 | 0 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
16 | 0010 | 0 | 0 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
17 | 0011 | 0 | 0 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
18 | 0012 | 0 | 0 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
19 | 0013 | 0 | 0 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
20 | 0014 | 0 | 0 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
21 | 0015 | 0 | 1 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
22 | 0016 | 0 | 1 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
23 | 0017 | 0 | 1 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
24 | 0018 | 0 | 1 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
25 | 0019 | 0 | 1 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
26 | 001A | 0 | 1 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
27 | 001B | 0 | 1 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
28 | 001C | 0 | 1 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
29 | 001D | 0 | 2 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
30 | 001E | 0 | 2 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
31 | 001F | 0 | 2 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
32 | 0020 | 0 | 2 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
33 | 0021 | 0 | 2 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
34 | 0022 | 0 | 2 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
35 | 0023 | 0 | 2 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
36 | 0024 | 0 | 2 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
37 | 0025 | 0 | 3 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
38 | 0026 | 0 | 3 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
39 | 0027 | 0 | 3 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
40 | 0028 | 0 | 3 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
41 | 0029 | 0 | 3 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
42 | 002A | 0 | 3 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
43 | 002B | 0 | 3 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
44 | 002C | 0 | 3 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
45 | 002D | 1 | 0 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
46 | 002E | 1 | 0 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
47 | 002F | 1 | 0 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
48 | 0030 | 1 | 0 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
49 | 0031 | 1 | 0 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
50 | 0032 | 1 | 0 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
51 | 0033 | 1 | 0 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
52 | 0034 | 1 | 0 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
53 | 0035 | 1 | 1 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
54 | 0036 | 1 | 1 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
55 | 0037 | 1 | 1 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
56 | 0038 | 1 | 1 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
57 | 0039 | 1 | 1 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
58 | 003A | 1 | 1 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
59 | 003B | 1 | 1 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
60 | 003C | 1 | 1 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
61 | 003D | 1 | 2 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
62 | 003E | 1 | 2 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
63 | 003F | 1 | 2 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
64 | 0040 | 1 | 2 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
65 | 0041 | 1 | 2 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
66 | 0042 | 1 | 2 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
67 | 0043 | 1 | 2 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
68 | 0044 | 1 | 2 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
69 | 0045 | 1 | 3 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
70 | 0046 | 1 | 3 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
71 | 0047 | 1 | 3 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
72 | 0048 | 1 | 3 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
73 | 0049 | 1 | 3 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
74 | 004A | 1 | 3 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
75 | 004B | 1 | 3 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
76 | 004C | 1 | 3 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
77 | 004D | 2 | 0 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
78 | 004E | 2 | 0 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
79 | 004F | 2 | 0 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
80 | 0050 | 2 | 0 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
81 | 0051 | 2 | 0 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
82 | 0052 | 2 | 0 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
83 | 0053 | 2 | 0 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
84 | 0054 | 2 | 0 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
85 | 0055 | 2 | 1 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
86 | 0056 | 2 | 1 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
87 | 0057 | 2 | 1 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
88 | 0058 | 2 | 1 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
89 | 0059 | 2 | 1 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
90 | 005A | 2 | 1 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
91 | 005B | 2 | 1 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
92 | 005C | 2 | 1 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
93 | 005D | 2 | 2 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
94 | 005E | 2 | 2 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
95 | 005F | 2 | 2 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
96 | 0060 | 2 | 2 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
97 | 0061 | 2 | 2 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
98 | 0062 | 2 | 2 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
99 | 0063 | 2 | 2 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
100 | 0064 | 2 | 2 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
101 | 0065 | 2 | 3 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
102 | 0066 | 2 | 3 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
103 | 0067 | 2 | 3 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
104 | 0068 | 2 | 3 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
105 | 0069 | 2 | 3 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
106 | 006A | 2 | 3 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
107 | 006B | 2 | 3 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
108 | 006C | 2 | 3 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
109 | 006D | 3 | 0 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
110 | 006E | 3 | 0 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
111 | 006F | 3 | 0 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
112 | 0070 | 3 | 0 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
113 | 0071 | 3 | 0 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
114 | 0072 | 3 | 0 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
115 | 0073 | 3 | 0 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
116 | 0074 | 3 | 0 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
117 | 0075 | 3 | 1 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
118 | 0076 | 3 | 1 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
119 | 0077 | 3 | 1 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
120 | 0078 | 3 | 1 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
121 | 0079 | 3 | 1 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
122 | 007A | 3 | 1 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
123 | 007B | 3 | 1 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
124 | 007C | 3 | 1 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
125 | 007D | 3 | 2 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
126 | 007E | 3 | 2 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
127 | 007F | 3 | 2 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
128 | 0080 | 3 | 2 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
129 | 0081 | 3 | 2 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
130 | 0082 | 3 | 2 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
131 | 0083 | 3 | 2 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
132 | 0084 | 3 | 2 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
133 | 0085 | 3 | 3 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
134 | 0086 | 3 | 3 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
135 | 0087 | 3 | 3 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
136 | 0088 | 3 | 3 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
137 | 0089 | 3 | 3 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
138 | 008A | 3 | 3 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
139 | 008B | 3 | 3 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
140 | 008C | 3 | 3 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
141 | 008D | N/A | N/A | N/A | Wilkinson rate feedback enables | Each bit turns on one ASIC's feedback enale for servoing VDLY | Bit 0: col 0, row 0 enable, Bit 1: col 0, row 1 enable, etc.
142 | 008E | N/A | N/A | N/A | Sampling rate feedback enables | Each bit turns on one ASIC's feedback enale for servoing VADJP/VADJN | Bit 0: col 0, row 0 enable, Bit 1: col 0, row 1 enable, etc.
143 | 008F | N/A | N/A | N/A | Trigger width feedback enables | Each bit turns on one ASIC's feedback enale for servoing WBIAS | Bit 0: col 0, row 0 enable, Bit 1: col 0, row 1 enable, etc.
144 | 0090 |  |  |  | Reserved (but not yet used)
145 | 0091 |  |  |  | Reserved (but not yet used)
146 | 0092 |  |  |  | Reserved (but not yet used)
147 | 0093 |  |  |  | Reserved (but not yet used)
148 | 0094 |  |  |  | Reserved (but not yet used)
149 | 0095 |  |  |  | Reserved (but not yet used)
150 | 0096 |  |  |  | Reserved (but not yet used)
151 | 0097 |  |  |  | Reserved (but not yet used)
152 | 0098 |  |  |  | Reserved (but not yet used)
153 | 0099 |  |  |  | Reserved (but not yet used)
154 | 009A |  |  |  | Reserved (but not yet used)
155 | 009B |  |  |  | Reserved (but not yet used)
156 | 009C |  |  |  | Reserved (but not yet used)
157 | 009D |  |  |  | Reserved (but not yet used)
158 | 009E |  |  |  | Reserved (but not yet used)
159 | 009F |  |  |  | Reserved (but not yet used)
160 | 00A0 |  |  |  | Reserved (but not yet used)
161 | 00A1 | G | G | G | First allowed analog storage window |  | Bits 8:0 store value, bits 16:9 unused
162 | 00A2 | G | G | G | Last allowed analog storage window |  | Bits 8:0 store value, bits 16:9 unused
163 | 00A3 | G | G | G | Maximum number of windows to look back |  | Bits 8:0 store value, bits 16:9 unused
164 | 00A4 | G | G | G | Minimum number of windows to look back |  | Bits 8:0 store value, bits 16:9 unused
165 | 00A5 | G | G | G | Software trigger and pedestal mode flags |  | Bits 8:0 - window to read out in pedestal mode, Bits 11:8 - unused, Bit 12: veto hardware triggers, Bit 13: veto software triggers, Bit 14: software trigger, Bit 15: flag to operate in pedestal mode
166 | 00A6 | N/A | N/A | N/A | Copy of the SCROD ID | Copy of the SCROD ID (read from EEPROM at startup, written here so the event builder can see it) | Bits 15:0, copy of SCROD ID - should be the decimal ID that the SCROD is labeled with
167 | 00A7 | N/A | N/A | N/A | Copy of the SCROD revision | Copy of the SCROD rev (read from EEPROM at startup, written here so the event builder can see it) | Bits 7:0, copy of the SCROD revision, Bits 15:8 reserved (e.g., SCROD revA2 should be 0x00A2)
168 | 00A8 | N/A | N/A | N/A | Event number to set [15:0] |  | LSB of event number to be set when set event register bit is toggled high
169 | 00A9 | N/A | N/A | N/A | Event number to set [31:16] |  | MSB of event number to be set when set event register bit is toggled high
170 | 00AA | N/A | N/A | N/A | Set event number |  | Bit 0: set event number on toggle high, Bits 15:1 unused
171 | 00AB | 0 | 0,1 | 7:0,7:0 | Force channel masks |  | Start from register 171 at col 0, row 0, ch 0, one bit per channel, increasing by channel, then row, then col
172 | 00AC | 0 | 2,3 | 7:0,7:0 | Force channel masks |  | See 171
173 | 00AD | 1 | 0,1 | 7:0,7:0 | Force channel masks |  | See 171
174 | 00AE | 1 | 2,3 | 7:0,7:0 | Force channel masks |  | See 171
175 | 00AF | 2 | 0,1 | 7:0,7:0 | Force channel masks |  | See 171
176 | 00B0 | 2 | 2,3 | 7:0,7:0 | Force channel masks |  | See 171
177 | 00B1 | 3 | 0,1 | 7:0,7:0 | Force channel masks |  | See 171
178 | 00B2 | 3 | 2,3 | 7:0,7:0 | Force channel masks |  | See 171
179 | 00B3 | 0 | 0,1 | 7:0,7:0 | Ignore channel masks |  | Start from register 179 at col 0, row 0, ch 0, one bit per channel, increasing by channel, then row, then col
180 | 00B4 | 0 | 2,3 | 7:0,7:0 | Ignore channel masks |  | See 179
181 | 00B5 | 1 | 0,1 | 7:0,7:0 | Ignore channel masks |  | See 179
182 | 00B6 | 1 | 2,3 | 7:0,7:0 | Ignore channel masks |  | See 179
183 | 00B7 | 2 | 0,1 | 7:0,7:0 | Ignore channel masks |  | See 179
184 | 00B8 | 2 | 2,3 | 7:0,7:0 | Ignore channel masks |  | See 179
185 | 00B9 | 3 | 0,1 | 7:0,7:0 | Ignore channel masks |  | See 179
186 | 00BA | 3 | 2,3 | 7:0,7:0 | Ignore channel masks |  | See 179
187 | 00BB | G | G | G | Region of interest adjust window |  | Bits 9:0 - a signed number that represents the offset for writing trigger bits to trigger memory. Use this to adjust if ROI is off by a bit. Bits 15:10 unused
188 | 00BC | G | G | G | Number of window pairs to sample after trigger |  | Bits 8:0 - unsigned number of window pairs (128 samples) to continue to sample after a trigger is received.
189 | 00BD |  |  |  | Reserved (but not yet used)
190 | 00BE |  |  |  | Reserved (but not yet used)
191 | 00BF |  |  |  | Reserved (but not yet used)
192 | 00C0 |  |  |  | Reserved (but not yet used)
193 | 00C1 |  |  |  | Reserved (but not yet used)
194 | 00C2 |  |  |  | Reserved (but not yet used)
195 | 00C3 |  |  |  | Reserved (but not yet used)
196 | 00C4 |  |  |  | Reserved (but not yet used)
197 | 00C5 |  |  |  | Reserved (but not yet used)
198 | 00C6 |  |  |  | Reserved (but not yet used)
199 | 00C7 |  |  |  | Reserved (but not yet used)
200 | 00C8 | G | G | G | DAC_BUF_BIAS_ISEL | Internal ASIC DAC buffer bias for driving ISEL - shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
201 | 00C9 | G | G | G | DAC_BUF_BIAS_VADJP | Internal ASIC DAC buffer bias for driving VadJP - shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
202 | 00CA | G | G | G | DAC_BUF_BIAS_VADJN | Internal ASIC DAC buffer bias for driving VadjN - shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
203 | 00CB | 0 | 0 | G | VBIAS | internal ASIC DAC bias for first stage transfer - unique to each ASIC to allow putting them into low power mode | Bits 15:12 unused, Bits 11:0 DAC value
204 | 00CC | 0 | 1 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
205 | 00CD | 0 | 2 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
206 | 00CE | 0 | 3 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
207 | 00CF | 1 | 0 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
208 | 00D0 | 1 | 1 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
209 | 00D1 | 1 | 2 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
210 | 00D2 | 1 | 3 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
211 | 00D3 | 2 | 0 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
212 | 00D4 | 2 | 1 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
213 | 00D5 | 2 | 2 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
214 | 00D6 | 2 | 3 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
215 | 00D7 | 3 | 0 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
216 | 00D8 | 3 | 1 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
217 | 00D9 | 3 | 2 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
218 | 00DA | 3 | 3 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
219 | 00DB | 0 | 0 | G | VBIAS2 | internal ASIC DAC bias for 2nd stage transfer - unique to each ASIC to allow putting them into low power mode | Bits 15:12 unused, Bits 11:0 DAC value
220 | 00DC | 0 | 1 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
221 | 00DD | 0 | 2 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
222 | 00DE | 0 | 3 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
223 | 00DF | 1 | 0 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
224 | 00E0 | 1 | 1 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
225 | 00E1 | 1 | 2 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
226 | 00E2 | 1 | 3 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
227 | 00E3 | 2 | 0 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
228 | 00E4 | 2 | 1 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
229 | 00E5 | 2 | 2 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
230 | 00E6 | 2 | 3 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
231 | 00E7 | 3 | 0 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
232 | 00E8 | 3 | 1 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
233 | 00E9 | 3 | 2 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
234 | 00EA | 3 | 3 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
235 | 00EB | 0 | 0 | G | WBIAS | Manual setting for trigger width bias - can be overridden by feedback (see register 143) | Bits 15:12 unused, Bits 11:0 DAC value
236 | 00EC | 0 | 1 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
237 | 00ED | 0 | 2 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
238 | 00EE | 0 | 3 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
239 | 00EF | 1 | 0 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
240 | 00F0 | 1 | 1 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
241 | 00F1 | 1 | 2 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
242 | 00F2 | 1 | 3 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
243 | 00F3 | 2 | 0 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
244 | 00F4 | 2 | 1 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
245 | 00F5 | 2 | 2 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
246 | 00F6 | 2 | 3 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
247 | 00F7 | 3 | 0 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
248 | 00F8 | 3 | 1 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
249 | 00F9 | 3 | 2 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
250 | 00FA | 3 | 3 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
251 | 00FB | 0 | 0 | G | VADJP | Manual setting for VADJP - can be overridden by feedback (see register 142) | Bits 15:12 unused, Bits 11:0 DAC value
252 | 00FC | 0 | 1 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
253 | 00FD | 0 | 2 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
254 | 00FE | 0 | 3 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
255 | 00FF | 1 | 0 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
256 | 0100 | 1 | 1 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
257 | 0101 | 1 | 2 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
258 | 0102 | 1 | 3 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
259 | 0103 | 2 | 0 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
260 | 0104 | 2 | 1 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
261 | 0105 | 2 | 2 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
262 | 0106 | 2 | 3 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
263 | 0107 | 3 | 0 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
264 | 0108 | 3 | 1 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
265 | 0109 | 3 | 2 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
266 | 010A | 3 | 3 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
267 | 010B | 0 | 0 | G | VADJN | Manual setting for VADJN - can be overridden by feedback (see register 142) | Bits 15:12 unused, Bits 11:0 DAC value
268 | 010C | 0 | 1 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
269 | 010D | 0 | 2 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
270 | 010E | 0 | 3 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
271 | 010F | 1 | 0 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
272 | 0110 | 1 | 1 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
273 | 0111 | 1 | 2 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
274 | 0112 | 1 | 3 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
275 | 0113 | 2 | 0 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
276 | 0114 | 2 | 1 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
277 | 0115 | 2 | 2 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
278 | 0116 | 2 | 3 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
279 | 0117 | 3 | 0 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
280 | 0118 | 3 | 1 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
281 | 0119 | 3 | 2 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
282 | 011A | 3 | 3 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
283 | 011B | 0 | 0 | G | VDLY | Manual setting for VDLY - wilkinson counter rate - can be overridden by feedback (see register 141) | Bits 15:12 unused, Bits 11:0 DAC value
284 | 011C | 0 | 1 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
285 | 011D | 0 | 2 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
286 | 011E | 0 | 3 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
287 | 011F | 1 | 0 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
288 | 0120 | 1 | 1 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
289 | 0121 | 1 | 2 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
290 | 0122 | 1 | 3 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
291 | 0123 | 2 | 0 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
292 | 0124 | 2 | 1 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
293 | 0125 | 2 | 2 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
294 | 0126 | 2 | 3 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
295 | 0127 | 3 | 0 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
296 | 0128 | 3 | 1 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
297 | 0129 | 3 | 2 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
298 | 012A | 3 | 3 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
299 | 012B | 0 | 0 | G | TIMING_SSP | Timing registers for SSP, lower bits are leading, upper bits are trailing | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
300 | 012C | 0 | 1 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
301 | 012D | 0 | 2 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
302 | 012E | 0 | 3 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
303 | 012F | 1 | 0 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
304 | 0130 | 1 | 1 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
305 | 0131 | 1 | 2 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
306 | 0132 | 1 | 3 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
307 | 0133 | 2 | 0 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
308 | 0134 | 2 | 1 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
309 | 0135 | 2 | 2 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
310 | 0136 | 2 | 3 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
311 | 0137 | 3 | 0 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
312 | 0138 | 3 | 1 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
313 | 0139 | 3 | 2 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
314 | 013A | 3 | 3 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
315 | 013B | 0 | 0 | G | TIMING_S1 | Timing registers for S1, lower bits are leading, upper bits are trailing | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
316 | 013C | 0 | 1 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
317 | 013D | 0 | 2 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
318 | 013E | 0 | 3 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
319 | 013F | 1 | 0 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
320 | 0140 | 1 | 1 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
321 | 0141 | 1 | 2 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
322 | 0142 | 1 | 3 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
323 | 0143 | 2 | 0 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
324 | 0144 | 2 | 1 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
325 | 0145 | 2 | 2 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
326 | 0146 | 2 | 3 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
327 | 0147 | 3 | 0 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
328 | 0148 | 3 | 1 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
329 | 0149 | 3 | 2 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
330 | 014A | 3 | 3 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
331 | 014B | 0 | 0 | G | TIMING_S2 | Timing registers for S2, lower bits are leading, upper bits are trailing | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
332 | 014C | 0 | 1 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
333 | 014D | 0 | 2 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
334 | 014E | 0 | 3 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
335 | 014F | 1 | 0 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
336 | 0150 | 1 | 1 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
337 | 0151 | 1 | 2 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
338 | 0152 | 1 | 3 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
339 | 0153 | 2 | 0 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
340 | 0154 | 2 | 1 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
341 | 0155 | 2 | 2 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
342 | 0156 | 2 | 3 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
343 | 0157 | 3 | 0 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
344 | 0158 | 3 | 1 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
345 | 0159 | 3 | 2 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
346 | 015A | 3 | 3 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
347 | 015B | 0 | 0 | G | TIMING_PHASE | Timing registers for PHASE, lower bits are leading, upper bits are trailing | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
348 | 015C | 0 | 1 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
349 | 015D | 0 | 2 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
350 | 015E | 0 | 3 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
351 | 015F | 1 | 0 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
352 | 0160 | 1 | 1 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
353 | 0161 | 1 | 2 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
354 | 0162 | 1 | 3 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
355 | 0163 | 2 | 0 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
356 | 0164 | 2 | 1 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
357 | 0165 | 2 | 2 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
358 | 0166 | 2 | 3 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
359 | 0167 | 3 | 0 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
360 | 0168 | 3 | 1 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
361 | 0169 | 3 | 2 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
362 | 016A | 3 | 3 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
363 | 016B | 0 | 0 | G | TIMING_WR_STRB | Timing registers for WR_STRB lower bits are leading, upper bits are trailing | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
364 | 016C | 0 | 1 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
365 | 016D | 0 | 2 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
366 | 016E | 0 | 3 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
367 | 016F | 1 | 0 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
368 | 0170 | 1 | 1 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
369 | 0171 | 1 | 2 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
370 | 0172 | 1 | 3 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
371 | 0173 | 2 | 0 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
372 | 0174 | 2 | 1 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
373 | 0175 | 2 | 2 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
374 | 0176 | 2 | 3 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
375 | 0177 | 3 | 0 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
376 | 0178 | 3 | 1 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
377 | 0179 | 3 | 2 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
378 | 017A | 3 | 3 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
379 | 017B | G | G | G | TIMING_REG & sampling sync bits | Register for internal ASIC timing signals (see sheet "IRS3B_TIMING_REGISTER") | Bits 7:0 timing reg (ASIC register), Bit 8: Signal for the sampling block to perform the sync - set to one after all taps/biases are programmed, Bits 10:9 Choose which phase to use among 4 PHAB phases
380 | 017C | N/A | N/A | N/A | MON_HEADER_MON_TIMING_ROW/COL (& RCOSSX) | Choose which MONTIMING and RCOSSX signal to send to the monitor headers | Bits 1:0 row select for MONTIMING/RCOSSX, bits 3:2 col select for MONTIMING/RCOSSX, bit 4: choose between MONTIMING and RCOSSX
381 | 017D |  |  |  | Reserved (but not yet used)
382 | 017E |  |  |  | Reserved (but not yet used)
383 | 017F |  |  |  | Reserved (but not yet used)
384 | 0180 | 0 | 0 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
385 | 0181 | 0 | 1 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
386 | 0182 | 0 | 2 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
387 | 0183 | 0 | 3 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
388 | 0184 | 1 | 0 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
389 | 0185 | 1 | 1 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
390 | 0186 | 1 | 2 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
391 | 0187 | 1 | 3 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
392 | 0188 | 2 | 0 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
393 | 0189 | 2 | 1 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
394 | 018A | 2 | 2 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
395 | 018B | 2 | 3 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
396 | 018C | 3 | 0 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
397 | 018D | 3 | 1 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
398 | 018E | 3 | 2 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
399 | 018F | 3 | 3 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
400 | 0190 | 0 | 0 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
401 | 0191 | 0 | 1 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
402 | 0192 | 0 | 2 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
403 | 0193 | 0 | 3 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
404 | 0194 | 1 | 0 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
405 | 0195 | 1 | 1 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
406 | 0196 | 1 | 2 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
407 | 0197 | 1 | 3 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
408 | 0198 | 2 | 0 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
409 | 0199 | 2 | 1 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
410 | 019A | 2 | 2 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
411 | 019B | 2 | 3 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
412 | 019C | 3 | 0 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
413 | 019D | 3 | 1 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
414 | 019E | 3 | 2 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
415 | 019F | 3 | 3 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
416 | 01A0 | G | G | G | USE_EXTERNAL_VADJ_DACS | LSB chooses whether to use the internal or external VadjP/N DACs | LSB only, others ignored
417 | 01A1 |  |  |  | Reserved (but not yet used)
418 | 01A2 |  |  |  | Reserved (but not yet used)
419 | 01A3 |  |  |  | Reserved (but not yet used)
420 | 01A4 |  |  |  | Reserved (but not yet used)
421 | 01A5 |  |  |  | Reserved (but not yet used)
422 | 01A6 |  |  |  | Reserved (but not yet used)
423 | 01A7 |  |  |  | Reserved (but not yet used)
424 | 01A8 |  |  |  | Reserved (but not yet used)
425 | 01A9 |  |  |  | Reserved (but not yet used)
426 | 01AA |  |  |  | Reserved (but not yet used)
427 | 01AB |  |  |  | Reserved (but not yet used)
428 | 01AC |  |  |  | Reserved (but not yet used)
429 | 01AD |  |  |  | Reserved (but not yet used)
430 | 01AE |  |  |  | Reserved (but not yet used)
431 | 01AF |  |  |  | Reserved (but not yet used)
432 | 01B0 |  |  |  | Reserved (but not yet used)
433 | 01B1 |  |  |  | Reserved (but not yet used)
434 | 01B2 |  |  |  | Reserved (but not yet used)
435 | 01B3 |  |  |  | Reserved (but not yet used)
436 | 01B4 |  |  |  | Reserved (but not yet used)
437 | 01B5 |  |  |  | Reserved (but not yet used)
438 | 01B6 |  |  |  | Reserved (but not yet used)
439 | 01B7 |  |  |  | Reserved (but not yet used)
440 | 01B8 |  |  |  | Reserved (but not yet used)
441 | 01B9 |  |  |  | Reserved (but not yet used)
442 | 01BA |  |  |  | Reserved (but not yet used)
443 | 01BB |  |  |  | Reserved (but not yet used)
444 | 01BC |  |  |  | Reserved (but not yet used)
445 | 01BD |  |  |  | Reserved (but not yet used)
446 | 01BE |  |  |  | Reserved (but not yet used)
447 | 01BF |  |  |  | Reserved (but not yet used)
448 | 01C0 |  |  |  | Reserved (but not yet used)
449 | 01C1 |  |  |  | Reserved (but not yet used)
450 | 01C2 |  |  |  | Reserved (but not yet used)
451 | 01C3 |  |  |  | Reserved (but not yet used)
452 | 01C4 |  |  |  | Reserved (but not yet used)
453 | 01C5 |  |  |  | Reserved (but not yet used)
454 | 01C6 |  |  |  | Reserved (but not yet used)
455 | 01C7 |  |  |  | Reserved (but not yet used)
456 | 01C8 |  |  |  | Reserved (but not yet used)
457 | 01C9 |  |  |  | Reserved (but not yet used)
458 | 01CA |  |  |  | Reserved (but not yet used)
459 | 01CB |  |  |  | Reserved (but not yet used)
460 | 01CC |  |  |  | Reserved (but not yet used)
461 | 01CD |  |  |  | Reserved (but not yet used)
462 | 01CE |  |  |  | Reserved (but not yet used)
463 | 01CF |  |  |  | Reserved (but not yet used)
464 | 01D0 |  |  |  | Reserved (but not yet used)
465 | 01D1 |  |  |  | Reserved (but not yet used)
466 | 01D2 |  |  |  | Reserved (but not yet used)
467 | 01D3 |  |  |  | Reserved (but not yet used)
468 | 01D4 |  |  |  | Reserved (but not yet used)
469 | 01D5 |  |  |  | Reserved (but not yet used)
470 | 01D6 |  |  |  | Reserved (but not yet used)
471 | 01D7 |  |  |  | Reserved (but not yet used)
472 | 01D8 |  |  |  | Reserved (but not yet used)
473 | 01D9 |  |  |  | Reserved (but not yet used)
474 | 01DA |  |  |  | Reserved (but not yet used)
475 | 01DB |  |  |  | Reserved (but not yet used)
476 | 01DC |  |  |  | Reserved (but not yet used)
477 | 01DD |  |  |  | Reserved (but not yet used)
478 | 01DE |  |  |  | Reserved (but not yet used)
479 | 01DF |  |  |  | Reserved (but not yet used)
480 | 01E0 |  |  |  | Reserved (but not yet used)
481 | 01E1 |  |  |  | Reserved (but not yet used)
482 | 01E2 |  |  |  | Reserved (but not yet used)
483 | 01E3 |  |  |  | Reserved (but not yet used)
484 | 01E4 |  |  |  | Reserved (but not yet used)
485 | 01E5 |  |  |  | Reserved (but not yet used)
486 | 01E6 |  |  |  | Reserved (but not yet used)
487 | 01E7 |  |  |  | Reserved (but not yet used)
488 | 01E8 |  |  |  | Reserved (but not yet used)
489 | 01E9 |  |  |  | Reserved (but not yet used)
490 | 01EA |  |  |  | Reserved (but not yet used)
491 | 01EB |  |  |  | Reserved (but not yet used)
492 | 01EC |  |  |  | Reserved (but not yet used)
493 | 01ED |  |  |  | Reserved (but not yet used)
494 | 01EE |  |  |  | Reserved (but not yet used)
495 | 01EF |  |  |  | Reserved (but not yet used)
496 | 01F0 |  |  |  | Reserved (but not yet used)
497 | 01F1 |  |  |  | Reserved (but not yet used)
498 | 01F2 |  |  |  | Reserved (but not yet used)
499 | 01F3 |  |  |  | Reserved (but not yet used)
500 | 01F4 | 0,1,2,3 | 0,1 | N/A | I2C_BUS[3]_Write | I2C write for row 0,1 temp sensors (x8), eeproms (x2), and gpios (x2) for cal signals (and SMPL_SEL_ANY) | Bits 15:13 unused, Bit 12: Send stop, Bit 11: Send acknowledge, Bit 10: Read byte, Bit 9: Send byte, Bit 8: Send start, Bit 7:0 Byte to send
501 | 01F5 | 0,1,2,3 | 2,3 | N/A | I2C_BUS[4]_Write | I2C write for row 2,3 temp sensors (x8), eeproms (x2), and gpios (x2) for cal signals (and SMPL_SEL_ANY) | Bits 15:13 unused, Bit 12: Send stop, Bit 11: Send acknowledge, Bit 10: Read byte, Bit 9: Send byte, Bit 8: Send start, Bit 7:0 Byte to send
502 | 01F6 | 0,1,2,3 | 0,1 | N/A | I2C_BUS[5]_Write | I2C write for row 0,1 GPIO serial for ASIC shift out signals | Bits 15:13 unused, Bit 12: Send stop, Bit 11: Send acknowledge, Bit 10: Read byte, Bit 9: Send byte, Bit 8: Send start, Bit 7:0 Byte to send
503 | 01F7 | 0,1,2,3 | 2,3 | N/A | I2C_BUS[6]_Write | I2C write for row 2,3 GPIO serial for ASIC shift out signals | Bits 15:13 unused, Bit 12: Send stop, Bit 11: Send acknowledge, Bit 10: Read byte, Bit 9: Send byte, Bit 8: Send start, Bit 7:0 Byte to send
504 | 01F8 | N/A | N/A | N/A | I2C_BUS[7]_Write | I2C write for interconnect rev C GPIO to control calibration signals | Bits 15:13 unused, Bit 12: Send stop, Bit 11: Send acknowledge, Bit 10: Read byte, Bit 9: Send byte, Bit 8: Send start, Bit 7:0 Byte to send
505 | 01F9 |  |  |  | Reserved (but not yet used)
506 | 01FA |  |  |  | Reserved (but not yet used)
507 | 01FB |  |  |  | Reserved (but not yet used)
508 | 01FC |  |  |  | Reserved (but not yet used)
509 | 01FD |  |  |  | Reserved (but not yet used)
510 | 01FE |  |  |  | Reserved (but not yet used)
511 | 01FF |  |  |  | Reserved (but not yet used)

## SCROD Read Registers
Register Number | (same in hex) | Col | Row | Ch | Name | Comments | Bit Mappings
0 | 0000 | N/A | N/A | N/A | LEDS | LEDs. Display SCROD ID on startup. FET must be installed properly. LED switch must be active. | One bit per LED
1 | 0001 | N/A | N/A | N/A | I2C_BUS[0]_Write | I2C write for SCROD EEPROM and temperature sensor on SCROD | Bits 15:13 unused, Bit 12: Send stop, Bit 11: Send acknowledge, Bit 10: Read byte, Bit 9: Send byte, Bit 8: Send start, Bit 7:0 Byte to send
2 | 0002 | N/A | N/A | N/A | I2C_BUS[1]_Write | I2C write for SCROD fiber transceiver 0 | Bits 15:13 unused, Bit 12: Send stop, Bit 11: Send acknowledge, Bit 10: Read byte, Bit 9: Send byte, Bit 8: Send start, Bit 7:0 Byte to send
3 | 0003 | N/A | N/A | N/A | I2C_BUS[2]_Write | I2C write for SCROD fiber transceiver 1 | Bits 15:13 unused, Bit 12: Send stop, Bit 11: Send acknowledge, Bit 10: Read byte, Bit 9: Send byte, Bit 8: Send start, Bit 7:0 Byte to send
4 | 0004 | N/A | N/A | N/A | Trigger interfaces | Trigger polarity, set column/row for scaler monitoring | Bit 15: trigger polarity (1 for rising edge, 0 for falling), Bits 14:4: unused, Bits 3:2: col select, Bits 1:0: row select (col/row selects choose which trigger scalers to read)
5 | 0005 | G | G | G | TRGBIAS | Internal ASIC trigger comparator bias, shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
6 | 0006 | G | G | G | TRGBIAS2 | Internal ASIC trigger comparator bias for reference channel, shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
7 | 0007 | G | G | G | TRGTHREF | Internal ASIC trigger threshold voltage for reference channel, shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
8 | 0008 | G | G | G | ISEL | Internal ASIC wilkinson ramp current source bias, shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
9 | 0009 | G | G | G | SBBIAS | Internal ASIC bias for the buffers that distribute the Wilkinson ramp, shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
10 | 000A | G | G | G | PUBIAS | Internal ASIC bias for pullup in wilkinson comparators, shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
11 | 000B | G | G | G | CMPBIAS | Internal ASIC bias for wilkinson comparators, shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
12 | 000C | G | G | G | internal_ASIC_DAC_BUF_BIASES | Internal ASIC biases for VBDbias, WBDbias, TCBbias, THDbias, Tbbias, TRGDbias, PDDbias, PUDbias, SBDbias, VDDbias, shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
13 | 000D | 0 | 0 | 0 | TRG_THRESH | Internal ASIC channel trigger thresholds - unique channel-to-channel | Bits 15:12 unused, Bits 11:0 DAC value
14 | 000E | 0 | 0 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
15 | 000F | 0 | 0 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
16 | 0010 | 0 | 0 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
17 | 0011 | 0 | 0 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
18 | 0012 | 0 | 0 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
19 | 0013 | 0 | 0 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
20 | 0014 | 0 | 0 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
21 | 0015 | 0 | 1 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
22 | 0016 | 0 | 1 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
23 | 0017 | 0 | 1 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
24 | 0018 | 0 | 1 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
25 | 0019 | 0 | 1 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
26 | 001A | 0 | 1 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
27 | 001B | 0 | 1 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
28 | 001C | 0 | 1 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
29 | 001D | 0 | 2 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
30 | 001E | 0 | 2 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
31 | 001F | 0 | 2 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
32 | 0020 | 0 | 2 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
33 | 0021 | 0 | 2 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
34 | 0022 | 0 | 2 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
35 | 0023 | 0 | 2 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
36 | 0024 | 0 | 2 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
37 | 0025 | 0 | 3 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
38 | 0026 | 0 | 3 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
39 | 0027 | 0 | 3 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
40 | 0028 | 0 | 3 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
41 | 0029 | 0 | 3 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
42 | 002A | 0 | 3 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
43 | 002B | 0 | 3 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
44 | 002C | 0 | 3 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
45 | 002D | 1 | 0 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
46 | 002E | 1 | 0 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
47 | 002F | 1 | 0 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
48 | 0030 | 1 | 0 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
49 | 0031 | 1 | 0 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
50 | 0032 | 1 | 0 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
51 | 0033 | 1 | 0 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
52 | 0034 | 1 | 0 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
53 | 0035 | 1 | 1 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
54 | 0036 | 1 | 1 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
55 | 0037 | 1 | 1 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
56 | 0038 | 1 | 1 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
57 | 0039 | 1 | 1 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
58 | 003A | 1 | 1 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
59 | 003B | 1 | 1 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
60 | 003C | 1 | 1 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
61 | 003D | 1 | 2 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
62 | 003E | 1 | 2 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
63 | 003F | 1 | 2 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
64 | 0040 | 1 | 2 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
65 | 0041 | 1 | 2 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
66 | 0042 | 1 | 2 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
67 | 0043 | 1 | 2 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
68 | 0044 | 1 | 2 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
69 | 0045 | 1 | 3 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
70 | 0046 | 1 | 3 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
71 | 0047 | 1 | 3 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
72 | 0048 | 1 | 3 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
73 | 0049 | 1 | 3 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
74 | 004A | 1 | 3 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
75 | 004B | 1 | 3 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
76 | 004C | 1 | 3 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
77 | 004D | 2 | 0 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
78 | 004E | 2 | 0 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
79 | 004F | 2 | 0 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
80 | 0050 | 2 | 0 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
81 | 0051 | 2 | 0 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
82 | 0052 | 2 | 0 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
83 | 0053 | 2 | 0 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
84 | 0054 | 2 | 0 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
85 | 0055 | 2 | 1 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
86 | 0056 | 2 | 1 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
87 | 0057 | 2 | 1 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
88 | 0058 | 2 | 1 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
89 | 0059 | 2 | 1 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
90 | 005A | 2 | 1 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
91 | 005B | 2 | 1 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
92 | 005C | 2 | 1 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
93 | 005D | 2 | 2 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
94 | 005E | 2 | 2 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
95 | 005F | 2 | 2 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
96 | 0060 | 2 | 2 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
97 | 0061 | 2 | 2 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
98 | 0062 | 2 | 2 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
99 | 0063 | 2 | 2 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
100 | 0064 | 2 | 2 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
101 | 0065 | 2 | 3 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
102 | 0066 | 2 | 3 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
103 | 0067 | 2 | 3 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
104 | 0068 | 2 | 3 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
105 | 0069 | 2 | 3 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
106 | 006A | 2 | 3 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
107 | 006B | 2 | 3 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
108 | 006C | 2 | 3 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
109 | 006D | 3 | 0 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
110 | 006E | 3 | 0 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
111 | 006F | 3 | 0 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
112 | 0070 | 3 | 0 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
113 | 0071 | 3 | 0 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
114 | 0072 | 3 | 0 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
115 | 0073 | 3 | 0 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
116 | 0074 | 3 | 0 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
117 | 0075 | 3 | 1 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
118 | 0076 | 3 | 1 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
119 | 0077 | 3 | 1 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
120 | 0078 | 3 | 1 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
121 | 0079 | 3 | 1 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
122 | 007A | 3 | 1 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
123 | 007B | 3 | 1 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
124 | 007C | 3 | 1 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
125 | 007D | 3 | 2 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
126 | 007E | 3 | 2 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
127 | 007F | 3 | 2 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
128 | 0080 | 3 | 2 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
129 | 0081 | 3 | 2 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
130 | 0082 | 3 | 2 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
131 | 0083 | 3 | 2 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
132 | 0084 | 3 | 2 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
133 | 0085 | 3 | 3 | 0 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
134 | 0086 | 3 | 3 | 1 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
135 | 0087 | 3 | 3 | 2 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
136 | 0088 | 3 | 3 | 3 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
137 | 0089 | 3 | 3 | 4 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
138 | 008A | 3 | 3 | 5 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
139 | 008B | 3 | 3 | 6 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
140 | 008C | 3 | 3 | 7 | TRG_THRESH | See reg 13 | Bits 15:12 unused, Bits 11:0 DAC value
141 | 008D | N/A | N/A | N/A | Wilkinson rate feedback enables | Each bit turns on one ASIC's feedback enale for servoing VDLY | Bit 0: col 0, row 0 enable, Bit 1: col 0, row 1 enable, etc.
142 | 008E | N/A | N/A | N/A | Sampling rate feedback enables | Each bit turns on one ASIC's feedback enale for servoing VADJP/VADJN | Bit 0: col 0, row 0 enable, Bit 1: col 0, row 1 enable, etc.
143 | 008F | N/A | N/A | N/A | Trigger width feedback enables | Each bit turns on one ASIC's feedback enale for servoing WBIAS | Bit 0: col 0, row 0 enable, Bit 1: col 0, row 1 enable, etc.
144 | 0090 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
145 | 0091 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
146 | 0092 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
147 | 0093 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
148 | 0094 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
149 | 0095 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
150 | 0096 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
151 | 0097 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
152 | 0098 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
153 | 0099 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
154 | 009A | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
155 | 009B | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
156 | 009C | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
157 | 009D | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
158 | 009E | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
159 | 009F | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
160 | 00A0 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
161 | 00A1 | G | G | G | First allowed analog storage window | 0 | Bits 8:0 store value, bits 16:9 unused
162 | 00A2 | G | G | G | Last allowed analog storage window | 0 | Bits 8:0 store value, bits 16:9 unused
163 | 00A3 | G | G | G | Maximum number of windows to look back | 0 | Bits 8:0 store value, bits 16:9 unused
164 | 00A4 | G | G | G | Minimum number of windows to look back | 0 | Bits 8:0 store value, bits 16:9 unused
165 | 00A5 | G | G | G | Software trigger and pedestal mode flags | 0 | Bits 8:0 - window to read out in pedestal mode, Bits 11:8 - unused, Bit 12: veto hardware triggers, Bit 13: veto software triggers, Bit 14: software trigger, Bit 15: flag to operate in pedestal mode
166 | 00A6 | N/A | N/A | N/A | Copy of the SCROD ID | Copy of the SCROD ID (read from EEPROM at startup, written here so the event builder can see it) | Bits 15:0, copy of SCROD ID - should be the decimal ID that the SCROD is labeled with
167 | 00A7 | N/A | N/A | N/A | Copy of the SCROD revision | Copy of the SCROD rev (read from EEPROM at startup, written here so the event builder can see it) | Bits 7:0, copy of the SCROD revision, Bits 15:8 reserved (e.g., SCROD revA2 should be 0x00A2)
168 | 00A8 | N/A | N/A | N/A | Event number to set [15:0] | 0 | LSB of event number to be set when set event register bit is toggled high
169 | 00A9 | N/A | N/A | N/A | Event number to set [31:16] | 0 | MSB of event number to be set when set event register bit is toggled high
170 | 00AA | N/A | N/A | N/A | Set event number | 0 | Bit 0: set event number on toggle high, Bits 15:1 unused
171 | 00AB | 0 | 0,1 | 7:0,7:0 | Force channel masks | 0 | Start from register 171 at col 0, row 0, ch 0, one bit per channel, increasing by channel, then row, then col
172 | 00AC | 0 | 2,3 | 7:0,7:0 | Force channel masks | 0 | See 171
173 | 00AD | 1 | 0,1 | 7:0,7:0 | Force channel masks | 0 | See 171
174 | 00AE | 1 | 2,3 | 7:0,7:0 | Force channel masks | 0 | See 171
175 | 00AF | 2 | 0,1 | 7:0,7:0 | Force channel masks | 0 | See 171
176 | 00B0 | 2 | 2,3 | 7:0,7:0 | Force channel masks | 0 | See 171
177 | 00B1 | 3 | 0,1 | 7:0,7:0 | Force channel masks | 0 | See 171
178 | 00B2 | 3 | 2,3 | 7:0,7:0 | Force channel masks | 0 | See 171
179 | 00B3 | 0 | 0,1 | 7:0,7:0 | Ignore channel masks | 0 | Start from register 179 at col 0, row 0, ch 0, one bit per channel, increasing by channel, then row, then col
180 | 00B4 | 0 | 2,3 | 7:0,7:0 | Ignore channel masks | 0 | See 179
181 | 00B5 | 1 | 0,1 | 7:0,7:0 | Ignore channel masks | 0 | See 179
182 | 00B6 | 1 | 2,3 | 7:0,7:0 | Ignore channel masks | 0 | See 179
183 | 00B7 | 2 | 0,1 | 7:0,7:0 | Ignore channel masks | 0 | See 179
184 | 00B8 | 2 | 2,3 | 7:0,7:0 | Ignore channel masks | 0 | See 179
185 | 00B9 | 3 | 0,1 | 7:0,7:0 | Ignore channel masks | 0 | See 179
186 | 00BA | 3 | 2,3 | 7:0,7:0 | Ignore channel masks | 0 | See 179
187 | 00BB | G | G | G | Region of interest adjust window | 0 | Bits 9:0 - a signed number that represents the offset for writing trigger bits to trigger memory. Use this to adjust if ROI is off by a bit. Bits 15:10 unused
188 | 00BC | G | G | G | Number of window pairs to sample after trigger | 0 | Bits 8:0 - unsigned number of window pairs (128 samples) to continue to sample after a trigger is received.
189 | 00BD | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
190 | 00BE | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
191 | 00BF | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
192 | 00C0 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
193 | 00C1 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
194 | 00C2 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
195 | 00C3 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
196 | 00C4 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
197 | 00C5 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
198 | 00C6 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
199 | 00C7 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
200 | 00C8 | G | G | G | DAC_BUF_BIAS_ISEL | Internal ASIC DAC buffer bias for driving ISEL - shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
201 | 00C9 | G | G | G | DAC_BUF_BIAS_VADJP | Internal ASIC DAC buffer bias for driving VadJP - shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
202 | 00CA | G | G | G | DAC_BUF_BIAS_VADJN | Internal ASIC DAC buffer bias for driving VadjN - shared for all ASICs | Bits 15:12 unused, Bits 11:0 DAC value
203 | 00CB | 0 | 0 | G | VBIAS | internal ASIC DAC bias for first stage transfer - unique to each ASIC to allow putting them into low power mode | Bits 15:12 unused, Bits 11:0 DAC value
204 | 00CC | 0 | 1 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
205 | 00CD | 0 | 2 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
206 | 00CE | 0 | 3 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
207 | 00CF | 1 | 0 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
208 | 00D0 | 1 | 1 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
209 | 00D1 | 1 | 2 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
210 | 00D2 | 1 | 3 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
211 | 00D3 | 2 | 0 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
212 | 00D4 | 2 | 1 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
213 | 00D5 | 2 | 2 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
214 | 00D6 | 2 | 3 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
215 | 00D7 | 3 | 0 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
216 | 00D8 | 3 | 1 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
217 | 00D9 | 3 | 2 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
218 | 00DA | 3 | 3 | G | VBIAS | see reg 203 | Bits 15:12 unused, Bits 11:0 DAC value
219 | 00DB | 0 | 0 | G | VBIAS2 | internal ASIC DAC bias for 2nd stage transfer - unique to each ASIC to allow putting them into low power mode | Bits 15:12 unused, Bits 11:0 DAC value
220 | 00DC | 0 | 1 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
221 | 00DD | 0 | 2 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
222 | 00DE | 0 | 3 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
223 | 00DF | 1 | 0 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
224 | 00E0 | 1 | 1 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
225 | 00E1 | 1 | 2 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
226 | 00E2 | 1 | 3 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
227 | 00E3 | 2 | 0 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
228 | 00E4 | 2 | 1 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
229 | 00E5 | 2 | 2 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
230 | 00E6 | 2 | 3 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
231 | 00E7 | 3 | 0 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
232 | 00E8 | 3 | 1 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
233 | 00E9 | 3 | 2 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
234 | 00EA | 3 | 3 | G | VBIAS2 | see reg 219 | Bits 15:12 unused, Bits 11:0 DAC value
235 | 00EB | 0 | 0 | G | WBIAS | Manual setting for trigger width bias - can be overridden by feedback (see register 143) | Bits 15:12 unused, Bits 11:0 DAC value
236 | 00EC | 0 | 1 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
237 | 00ED | 0 | 2 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
238 | 00EE | 0 | 3 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
239 | 00EF | 1 | 0 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
240 | 00F0 | 1 | 1 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
241 | 00F1 | 1 | 2 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
242 | 00F2 | 1 | 3 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
243 | 00F3 | 2 | 0 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
244 | 00F4 | 2 | 1 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
245 | 00F5 | 2 | 2 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
246 | 00F6 | 2 | 3 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
247 | 00F7 | 3 | 0 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
248 | 00F8 | 3 | 1 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
249 | 00F9 | 3 | 2 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
250 | 00FA | 3 | 3 | G | WBIAS | see reg 235 | Bits 15:12 unused, Bits 11:0 DAC value
251 | 00FB | 0 | 0 | G | VADJP | Manual setting for VADJP - can be overridden by feedback (see register 142) | Bits 15:12 unused, Bits 11:0 DAC value
252 | 00FC | 0 | 1 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
253 | 00FD | 0 | 2 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
254 | 00FE | 0 | 3 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
255 | 00FF | 1 | 0 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
256 | 0100 | 1 | 1 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
257 | 0101 | 1 | 2 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
258 | 0102 | 1 | 3 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
259 | 0103 | 2 | 0 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
260 | 0104 | 2 | 1 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
261 | 0105 | 2 | 2 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
262 | 0106 | 2 | 3 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
263 | 0107 | 3 | 0 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
264 | 0108 | 3 | 1 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
265 | 0109 | 3 | 2 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
266 | 010A | 3 | 3 | G | VADJP | see reg 251 | Bits 15:12 unused, Bits 11:0 DAC value
267 | 010B | 0 | 0 | G | VADJN | Manual setting for VADJN - can be overridden by feedback (see register 142) | Bits 15:12 unused, Bits 11:0 DAC value
268 | 010C | 0 | 1 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
269 | 010D | 0 | 2 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
270 | 010E | 0 | 3 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
271 | 010F | 1 | 0 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
272 | 0110 | 1 | 1 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
273 | 0111 | 1 | 2 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
274 | 0112 | 1 | 3 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
275 | 0113 | 2 | 0 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
276 | 0114 | 2 | 1 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
277 | 0115 | 2 | 2 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
278 | 0116 | 2 | 3 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
279 | 0117 | 3 | 0 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
280 | 0118 | 3 | 1 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
281 | 0119 | 3 | 2 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
282 | 011A | 3 | 3 | G | VADJN | see reg 267 | Bits 15:12 unused, Bits 11:0 DAC value
283 | 011B | 0 | 0 | G | VDLY | Manual setting for VDLY - wilkinson counter rate - can be overridden by feedback (see register 141) | Bits 15:12 unused, Bits 11:0 DAC value
284 | 011C | 0 | 1 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
285 | 011D | 0 | 2 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
286 | 011E | 0 | 3 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
287 | 011F | 1 | 0 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
288 | 0120 | 1 | 1 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
289 | 0121 | 1 | 2 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
290 | 0122 | 1 | 3 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
291 | 0123 | 2 | 0 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
292 | 0124 | 2 | 1 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
293 | 0125 | 2 | 2 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
294 | 0126 | 2 | 3 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
295 | 0127 | 3 | 0 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
296 | 0128 | 3 | 1 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
297 | 0129 | 3 | 2 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
298 | 012A | 3 | 3 | G | VDLY | see reg 283 | Bits 15:12 unused, Bits 11:0 DAC value
299 | 012B | 0 | 0 | G | TIMING_SSP | Timing registers for SSP, lower bits are leading, upper bits are trailing | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
300 | 012C | 0 | 1 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
301 | 012D | 0 | 2 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
302 | 012E | 0 | 3 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
303 | 012F | 1 | 0 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
304 | 0130 | 1 | 1 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
305 | 0131 | 1 | 2 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
306 | 0132 | 1 | 3 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
307 | 0133 | 2 | 0 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
308 | 0134 | 2 | 1 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
309 | 0135 | 2 | 2 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
310 | 0136 | 2 | 3 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
311 | 0137 | 3 | 0 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
312 | 0138 | 3 | 1 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
313 | 0139 | 3 | 2 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
314 | 013A | 3 | 3 | G | TIMING_SSP | see reg 299 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
315 | 013B | 0 | 0 | G | TIMING_S1 | Timing registers for S1, lower bits are leading, upper bits are trailing | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
316 | 013C | 0 | 1 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
317 | 013D | 0 | 2 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
318 | 013E | 0 | 3 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
319 | 013F | 1 | 0 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
320 | 0140 | 1 | 1 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
321 | 0141 | 1 | 2 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
322 | 0142 | 1 | 3 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
323 | 0143 | 2 | 0 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
324 | 0144 | 2 | 1 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
325 | 0145 | 2 | 2 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
326 | 0146 | 2 | 3 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
327 | 0147 | 3 | 0 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
328 | 0148 | 3 | 1 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
329 | 0149 | 3 | 2 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
330 | 014A | 3 | 3 | G | TIMING_S1 | see reg 315 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
331 | 014B | 0 | 0 | G | TIMING_S2 | Timing registers for S2, lower bits are leading, upper bits are trailing | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
332 | 014C | 0 | 1 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
333 | 014D | 0 | 2 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
334 | 014E | 0 | 3 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
335 | 014F | 1 | 0 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
336 | 0150 | 1 | 1 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
337 | 0151 | 1 | 2 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
338 | 0152 | 1 | 3 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
339 | 0153 | 2 | 0 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
340 | 0154 | 2 | 1 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
341 | 0155 | 2 | 2 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
342 | 0156 | 2 | 3 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
343 | 0157 | 3 | 0 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
344 | 0158 | 3 | 1 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
345 | 0159 | 3 | 2 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
346 | 015A | 3 | 3 | G | TIMING_S2 | see reg 331 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
347 | 015B | 0 | 0 | G | TIMING_PHASE | Timing registers for PHASE, lower bits are leading, upper bits are trailing | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
348 | 015C | 0 | 1 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
349 | 015D | 0 | 2 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
350 | 015E | 0 | 3 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
351 | 015F | 1 | 0 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
352 | 0160 | 1 | 1 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
353 | 0161 | 1 | 2 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
354 | 0162 | 1 | 3 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
355 | 0163 | 2 | 0 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
356 | 0164 | 2 | 1 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
357 | 0165 | 2 | 2 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
358 | 0166 | 2 | 3 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
359 | 0167 | 3 | 0 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
360 | 0168 | 3 | 1 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
361 | 0169 | 3 | 2 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
362 | 016A | 3 | 3 | G | TIMING_PHASE | see reg 347 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
363 | 016B | 0 | 0 | G | TIMING_WR_STRB | Timing registers for WR_STRB lower bits are leading, upper bits are trailing | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
364 | 016C | 0 | 1 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
365 | 016D | 0 | 2 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
366 | 016E | 0 | 3 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
367 | 016F | 1 | 0 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
368 | 0170 | 1 | 1 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
369 | 0171 | 1 | 2 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
370 | 0172 | 1 | 3 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
371 | 0173 | 2 | 0 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
372 | 0174 | 2 | 1 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
373 | 0175 | 2 | 2 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
374 | 0176 | 2 | 3 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
375 | 0177 | 3 | 0 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
376 | 0178 | 3 | 1 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
377 | 0179 | 3 | 2 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
378 | 017A | 3 | 3 | G | TIMING_WR_STRB | see reg 365 | Bits 7:0 set the timing of the leading edge, bits 15:8 set the timing of the trailing edge
379 | 017B | G | G | G | TIMING_REG & sampling sync bits | Register for internal ASIC timing signals (see sheet "IRS3B_TIMING_REGISTER") | Bits 7:0 timing reg (ASIC register), Bit 8: Signal for the sampling block to perform the sync - set to one after all taps/biases are programmed, Bits 10:9 Choose which phase to use among 4 PHAB phases
380 | 017C | N/A | N/A | N/A | MON_HEADER_MON_TIMING_ROW/COL (& RCOSSX) | Choose which MONTIMING and RCOSSX signal to send to the monitor headers | Bits 1:0 row select for MONTIMING/RCOSSX, bits 3:2 col select for MONTIMING/RCOSSX, bit 4: choose between MONTIMING and RCOSSX
381 | 017D | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
382 | 017E | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
383 | 017F | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
384 | 0180 | 0 | 0 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
385 | 0181 | 0 | 1 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
386 | 0182 | 0 | 2 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
387 | 0183 | 0 | 3 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
388 | 0184 | 1 | 0 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
389 | 0185 | 1 | 1 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
390 | 0186 | 1 | 2 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
391 | 0187 | 1 | 3 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
392 | 0188 | 2 | 0 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
393 | 0189 | 2 | 1 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
394 | 018A | 2 | 2 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
395 | 018B | 2 | 3 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
396 | 018C | 3 | 0 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
397 | 018D | 3 | 1 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
398 | 018E | 3 | 2 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
399 | 018F | 3 | 3 | G | WILKINSON_TARGET | Target count rate value for Wilkinson feedback loop | 16 bit unsigned counter
400 | 0190 | 0 | 0 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
401 | 0191 | 0 | 1 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
402 | 0192 | 0 | 2 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
403 | 0193 | 0 | 3 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
404 | 0194 | 1 | 0 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
405 | 0195 | 1 | 1 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
406 | 0196 | 1 | 2 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
407 | 0197 | 1 | 3 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
408 | 0198 | 2 | 0 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
409 | 0199 | 2 | 1 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
410 | 019A | 2 | 2 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
411 | 019B | 2 | 3 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
412 | 019C | 3 | 0 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
413 | 019D | 3 | 1 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
414 | 019E | 3 | 2 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
415 | 019F | 3 | 3 | G | RCO_TARGET | Target count rate value for RCO-based sampling rate feedback loop | 16 bit unsigned counter
416 | 01A0 | G | G | G | USE_EXTERNAL_VADJ_DACS | LSB chooses whether to use the internal or external VadjP/N DACs | LSB only, others ignored
417 | 01A1 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
418 | 01A2 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
419 | 01A3 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
420 | 01A4 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
421 | 01A5 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
422 | 01A6 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
423 | 01A7 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
424 | 01A8 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
425 | 01A9 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
426 | 01AA | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
427 | 01AB | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
428 | 01AC | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
429 | 01AD | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
430 | 01AE | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
431 | 01AF | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
432 | 01B0 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
433 | 01B1 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
434 | 01B2 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
435 | 01B3 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
436 | 01B4 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
437 | 01B5 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
438 | 01B6 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
439 | 01B7 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
440 | 01B8 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
441 | 01B9 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
442 | 01BA | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
443 | 01BB | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
444 | 01BC | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
445 | 01BD | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
446 | 01BE | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
447 | 01BF | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
448 | 01C0 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
449 | 01C1 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
450 | 01C2 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
451 | 01C3 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
452 | 01C4 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
453 | 01C5 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
454 | 01C6 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
455 | 01C7 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
456 | 01C8 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
457 | 01C9 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
458 | 01CA | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
459 | 01CB | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
460 | 01CC | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
461 | 01CD | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
462 | 01CE | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
463 | 01CF | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
464 | 01D0 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
465 | 01D1 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
466 | 01D2 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
467 | 01D3 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
468 | 01D4 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
469 | 01D5 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
470 | 01D6 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
471 | 01D7 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
472 | 01D8 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
473 | 01D9 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
474 | 01DA | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
475 | 01DB | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
476 | 01DC | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
477 | 01DD | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
478 | 01DE | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
479 | 01DF | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
480 | 01E0 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
481 | 01E1 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
482 | 01E2 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
483 | 01E3 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
484 | 01E4 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
485 | 01E5 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
486 | 01E6 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
487 | 01E7 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
488 | 01E8 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
489 | 01E9 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
490 | 01EA | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
491 | 01EB | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
492 | 01EC | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
493 | 01ED | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
494 | 01EE | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
495 | 01EF | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
496 | 01F0 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
497 | 01F1 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
498 | 01F2 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
499 | 01F3 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
500 | 01F4 | 0,1,2,3 | 0,1 | N/A | I2C_BUS[3]_Write | I2C write for row 0,1 temp sensors (x8), eeproms (x2), and gpios (x2) for cal signals (and SMPL_SEL_ANY) | Bits 15:13 unused, Bit 12: Send stop, Bit 11: Send acknowledge, Bit 10: Read byte, Bit 9: Send byte, Bit 8: Send start, Bit 7:0 Byte to send
501 | 01F5 | 0,1,2,3 | 2,3 | N/A | I2C_BUS[4]_Write | I2C write for row 2,3 temp sensors (x8), eeproms (x2), and gpios (x2) for cal signals (and SMPL_SEL_ANY) | Bits 15:13 unused, Bit 12: Send stop, Bit 11: Send acknowledge, Bit 10: Read byte, Bit 9: Send byte, Bit 8: Send start, Bit 7:0 Byte to send
502 | 01F6 | 0,1,2,3 | 0,1 | N/A | I2C_BUS[5]_Write | I2C write for row 0,1 GPIO serial for ASIC shift out signals | Bits 15:13 unused, Bit 12: Send stop, Bit 11: Send acknowledge, Bit 10: Read byte, Bit 9: Send byte, Bit 8: Send start, Bit 7:0 Byte to send
503 | 01F7 | 0,1,2,3 | 2,3 | N/A | I2C_BUS[6]_Write | I2C write for row 2,3 GPIO serial for ASIC shift out signals | Bits 15:13 unused, Bit 12: Send stop, Bit 11: Send acknowledge, Bit 10: Read byte, Bit 9: Send byte, Bit 8: Send start, Bit 7:0 Byte to send
504 | 01F8 | N/A | N/A | N/A | I2C_BUS[7]_Write | I2C write for interconnect rev C GPIO to control calibration signals | Bits 15:13 unused, Bit 12: Send stop, Bit 11: Send acknowledge, Bit 10: Read byte, Bit 9: Send byte, Bit 8: Send start, Bit 7:0 Byte to send
505 | 01F9 | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
506 | 01FA | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
507 | 01FB | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
508 | 01FC | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
509 | 01FD | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
510 | 01FE | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
511 | 01FF | 0 | 0 | 0 | Reserved (but not yet used) | 0 | 0
512 | 0200 | N/A | N/A | N/A | I2C_BUS[0]_Read | I2C read for SCROD EEPROM and temperature sensor on SCROD | Bit 15: I2C acknowledged, Bit 14: I2C busy, Bits 13:8: unused, Bit 7:0: byte received
513 | 0201 | N/A | N/A | N/A | I2C_BUS[1]_Read | I2C read for SCROD fiber transceiver 0 | Bit 15: I2C acknowledged, Bit 14: I2C busy, Bits 13:8: unused, Bit 7:0: byte received
514 | 0202 | N/A | N/A | N/A | I2C_BUS[2]_Read | I2C read for SCROD fiber transceiver 1 | Bit 15: I2C acknowledged, Bit 14: I2C busy, Bits 13:8: unused, Bit 7:0: byte received
515 | 0203 | N/A | N/A | N/A | I2C_BUS[3]_Read | I2C read for row 0,1 temp sensors (x8), eeproms (x2), and gpios (x2) for cal signals (and SMPL_SEL_ANY) | Bit 15: I2C acknowledged, Bit 14: I2C busy, Bits 13:8: unused, Bit 7:0: byte received
516 | 0204 | N/A | N/A | N/A | I2C_BUS[4]_Read | I2C read for row 2,3 temp sensors (x8), eeproms (x2), and gpios (x2) for cal signals (and SMPL_SEL_ANY) | Bit 15: I2C acknowledged, Bit 14: I2C busy, Bits 13:8: unused, Bit 7:0: byte received
517 | 0205 | N/A | N/A | N/A | I2C_BUS[5]_Read | I2C read for row 0,1 GPIO serial for ASIC shift out signals | Bit 15: I2C acknowledged, Bit 14: I2C busy, Bits 13:8: unused, Bit 7:0: byte received
518 | 0206 | N/A | N/A | N/A | I2C_BUS[6]_Read | I2C read for row 2,3 GPIO serial for ASIC shift out signals | Bit 15: I2C acknowledged, Bit 14: I2C busy, Bits 13:8: unused, Bit 7:0: byte received
519 | 0207 | N/A | N/A | N/A | I2C_BUS[7]_Read | I2C read for interconnect rev C GPIO to control calibration signals | Bit 15: I2C acknowledged, Bit 14: I2C busy, Bits 13:8: unused, Bit 7:0: byte received
520 | 0208 | See comments |  | 0 | ASIC scaler value | Scalers for ASIC as selected by write register 4 | 16 bit unsigned number
521 | 0209 | See comments |  | 1 | ASIC scaler value | Scalers for ASIC as selected by write register 4 | 16 bit unsigned number
522 | 020A | See comments |  | 2 | ASIC scaler value | Scalers for ASIC as selected by write register 4 | 16 bit unsigned number
523 | 020B | See comments |  | 3 | ASIC scaler value | Scalers for ASIC as selected by write register 4 | 16 bit unsigned number
524 | 020C | See comments |  | 4 | ASIC scaler value | Scalers for ASIC as selected by write register 4 | 16 bit unsigned number
525 | 020D | See comments |  | 5 | ASIC scaler value | Scalers for ASIC as selected by write register 4 | 16 bit unsigned number
526 | 020E | See comments |  | 6 | ASIC scaler value | Scalers for ASIC as selected by write register 4 | 16 bit unsigned number
527 | 020F | See comments |  | 7 | ASIC scaler value | Scalers for ASIC as selected by write register 4 | 16 bit unsigned number
528 | 0210 | 0 | 0 | N/A | WILK_COUNTER | Counter to monitor the Wilkinson clock rate for a given ASIC | 16 bit unsigned counter
529 | 0211 | 0 | 1 | N/A | WILK_COUNTER | see reg 528 | 16 bit unsigned counter
530 | 0212 | 0 | 2 | N/A | WILK_COUNTER | see reg 528 | 16 bit unsigned counter
531 | 0213 | 0 | 3 | N/A | WILK_COUNTER | see reg 528 | 16 bit unsigned counter
532 | 0214 | 1 | 0 | N/A | WILK_COUNTER | see reg 528 | 16 bit unsigned counter
533 | 0215 | 1 | 1 | N/A | WILK_COUNTER | see reg 528 | 16 bit unsigned counter
534 | 0216 | 1 | 2 | N/A | WILK_COUNTER | see reg 528 | 16 bit unsigned counter
535 | 0217 | 1 | 3 | N/A | WILK_COUNTER | see reg 528 | 16 bit unsigned counter
536 | 0218 | 2 | 0 | N/A | WILK_COUNTER | see reg 528 | 16 bit unsigned counter
537 | 0219 | 2 | 1 | N/A | WILK_COUNTER | see reg 528 | 16 bit unsigned counter
538 | 021A | 2 | 2 | N/A | WILK_COUNTER | see reg 528 | 16 bit unsigned counter
539 | 021B | 2 | 3 | N/A | WILK_COUNTER | see reg 528 | 16 bit unsigned counter
540 | 021C | 3 | 0 | N/A | WILK_COUNTER | see reg 528 | 16 bit unsigned counter
541 | 021D | 3 | 1 | N/A | WILK_COUNTER | see reg 528 | 16 bit unsigned counter
542 | 021E | 3 | 2 | N/A | WILK_COUNTER | see reg 528 | 16 bit unsigned counter
543 | 021F | 3 | 3 | N/A | WILK_COUNTER | see reg 528 | 16 bit unsigned counter
544 | 0220 | 0 | 0 | N/A | VDLY_FB | Feedback loop desired setting of VDLY | Bits 15:12 unused, Bits 11:0 DAC value
545 | 0221 | 0 | 1 | N/A | VDLY_FB | see reg 544 | Bits 15:12 unused, Bits 11:0 DAC value
546 | 0222 | 0 | 2 | N/A | VDLY_FB | see reg 544 | Bits 15:12 unused, Bits 11:0 DAC value
547 | 0223 | 0 | 3 | N/A | VDLY_FB | see reg 544 | Bits 15:12 unused, Bits 11:0 DAC value
548 | 0224 | 1 | 0 | N/A | VDLY_FB | see reg 544 | Bits 15:12 unused, Bits 11:0 DAC value
549 | 0225 | 1 | 1 | N/A | VDLY_FB | see reg 544 | Bits 15:12 unused, Bits 11:0 DAC value
550 | 0226 | 1 | 2 | N/A | VDLY_FB | see reg 544 | Bits 15:12 unused, Bits 11:0 DAC value
551 | 0227 | 1 | 3 | N/A | VDLY_FB | see reg 544 | Bits 15:12 unused, Bits 11:0 DAC value
552 | 0228 | 2 | 0 | N/A | VDLY_FB | see reg 544 | Bits 15:12 unused, Bits 11:0 DAC value
553 | 0229 | 2 | 1 | N/A | VDLY_FB | see reg 544 | Bits 15:12 unused, Bits 11:0 DAC value
554 | 022A | 2 | 2 | N/A | VDLY_FB | see reg 544 | Bits 15:12 unused, Bits 11:0 DAC value
555 | 022B | 2 | 3 | N/A | VDLY_FB | see reg 544 | Bits 15:12 unused, Bits 11:0 DAC value
556 | 022C | 3 | 0 | N/A | VDLY_FB | see reg 544 | Bits 15:12 unused, Bits 11:0 DAC value
557 | 022D | 3 | 1 | N/A | VDLY_FB | see reg 544 | Bits 15:12 unused, Bits 11:0 DAC value
558 | 022E | 3 | 2 | N/A | VDLY_FB | see reg 544 | Bits 15:12 unused, Bits 11:0 DAC value
559 | 022F | 3 | 3 | N/A | VDLY_FB | see reg 544 | Bits 15:12 unused, Bits 11:0 DAC value
560 | 0230 | 0 | 0 | N/A | RCO_COUNTER | SST RCO counter to monitor sampling rate | 16 bit unsigned counter
561 | 0231 | 0 | 1 | N/A | RCO_COUNTER | see reg 560 | 16 bit unsigned counter
562 | 0232 | 0 | 2 | N/A | RCO_COUNTER | see reg 560 | 16 bit unsigned counter
563 | 0233 | 0 | 3 | N/A | RCO_COUNTER | see reg 560 | 16 bit unsigned counter
564 | 0234 | 1 | 0 | N/A | RCO_COUNTER | see reg 560 | 16 bit unsigned counter
565 | 0235 | 1 | 1 | N/A | RCO_COUNTER | see reg 560 | 16 bit unsigned counter
566 | 0236 | 1 | 2 | N/A | RCO_COUNTER | see reg 560 | 16 bit unsigned counter
567 | 0237 | 1 | 3 | N/A | RCO_COUNTER | see reg 560 | 16 bit unsigned counter
568 | 0238 | 2 | 0 | N/A | RCO_COUNTER | see reg 560 | 16 bit unsigned counter
569 | 0239 | 2 | 1 | N/A | RCO_COUNTER | see reg 560 | 16 bit unsigned counter
570 | 023A | 2 | 2 | N/A | RCO_COUNTER | see reg 560 | 16 bit unsigned counter
571 | 023B | 2 | 3 | N/A | RCO_COUNTER | see reg 560 | 16 bit unsigned counter
572 | 023C | 3 | 0 | N/A | RCO_COUNTER | see reg 560 | 16 bit unsigned counter
573 | 023D | 3 | 1 | N/A | RCO_COUNTER | see reg 560 | 16 bit unsigned counter
574 | 023E | 3 | 2 | N/A | RCO_COUNTER | see reg 560 | 16 bit unsigned counter
575 | 023F | 3 | 3 | N/A | RCO_COUNTER | see reg 560 | 16 bit unsigned counter
576 | 0240 | 0 | 0 | N/A | VADJN_FB | Feedback DAC value of VADJN | Bits 15:12 unused, Bits 11:0 DAC value
577 | 0241 | 0 | 1 | N/A | VADJN_FB | see reg 576 | Bits 15:12 unused, Bits 11:0 DAC value
578 | 0242 | 0 | 2 | N/A | VADJN_FB | see reg 576 | Bits 15:12 unused, Bits 11:0 DAC value
579 | 0243 | 0 | 3 | N/A | VADJN_FB | see reg 576 | Bits 15:12 unused, Bits 11:0 DAC value
580 | 0244 | 1 | 0 | N/A | VADJN_FB | see reg 576 | Bits 15:12 unused, Bits 11:0 DAC value
581 | 0245 | 1 | 1 | N/A | VADJN_FB | see reg 576 | Bits 15:12 unused, Bits 11:0 DAC value
582 | 0246 | 1 | 2 | N/A | VADJN_FB | see reg 576 | Bits 15:12 unused, Bits 11:0 DAC value
583 | 0247 | 1 | 3 | N/A | VADJN_FB | see reg 576 | Bits 15:12 unused, Bits 11:0 DAC value
584 | 0248 | 2 | 0 | N/A | VADJN_FB | see reg 576 | Bits 15:12 unused, Bits 11:0 DAC value
585 | 0249 | 2 | 1 | N/A | VADJN_FB | see reg 576 | Bits 15:12 unused, Bits 11:0 DAC value
586 | 024A | 2 | 2 | N/A | VADJN_FB | see reg 576 | Bits 15:12 unused, Bits 11:0 DAC value
587 | 024B | 2 | 3 | N/A | VADJN_FB | see reg 576 | Bits 15:12 unused, Bits 11:0 DAC value
588 | 024C | 3 | 0 | N/A | VADJN_FB | see reg 576 | Bits 15:12 unused, Bits 11:0 DAC value
589 | 024D | 3 | 1 | N/A | VADJN_FB | see reg 576 | Bits 15:12 unused, Bits 11:0 DAC value
590 | 024E | 3 | 2 | N/A | VADJN_FB | see reg 576 | Bits 15:12 unused, Bits 11:0 DAC value
591 | 024F | 3 | 3 | N/A | VADJN_FB | see reg 576 | Bits 15:12 unused, Bits 11:0 DAC value
592 | 0250 | 0 | 0 | N/A | VADJP_FB | Feedback DAC value of VADJN | Bits 15:12 unused, Bits 11:0 DAC value
593 | 0251 | 0 | 1 | N/A | VADJP_FB | see reg 592 | Bits 15:12 unused, Bits 11:0 DAC value
594 | 0252 | 0 | 2 | N/A | VADJP_FB | see reg 592 | Bits 15:12 unused, Bits 11:0 DAC value
595 | 0253 | 0 | 3 | N/A | VADJP_FB | see reg 592 | Bits 15:12 unused, Bits 11:0 DAC value
596 | 0254 | 1 | 0 | N/A | VADJP_FB | see reg 592 | Bits 15:12 unused, Bits 11:0 DAC value
597 | 0255 | 1 | 1 | N/A | VADJP_FB | see reg 592 | Bits 15:12 unused, Bits 11:0 DAC value
598 | 0256 | 1 | 2 | N/A | VADJP_FB | see reg 592 | Bits 15:12 unused, Bits 11:0 DAC value
599 | 0257 | 1 | 3 | N/A | VADJP_FB | see reg 592 | Bits 15:12 unused, Bits 11:0 DAC value
600 | 0258 | 2 | 0 | N/A | VADJP_FB | see reg 592 | Bits 15:12 unused, Bits 11:0 DAC value
601 | 0259 | 2 | 1 | N/A | VADJP_FB | see reg 592 | Bits 15:12 unused, Bits 11:0 DAC value
602 | 025A | 2 | 2 | N/A | VADJP_FB | see reg 592 | Bits 15:12 unused, Bits 11:0 DAC value
603 | 025B | 2 | 3 | N/A | VADJP_FB | see reg 592 | Bits 15:12 unused, Bits 11:0 DAC value
604 | 025C | 3 | 0 | N/A | VADJP_FB | see reg 592 | Bits 15:12 unused, Bits 11:0 DAC value
605 | 025D | 3 | 1 | N/A | VADJP_FB | see reg 592 | Bits 15:12 unused, Bits 11:0 DAC value
606 | 025E | 3 | 2 | N/A | VADJP_FB | see reg 592 | Bits 15:12 unused, Bits 11:0 DAC value
607 | 025F | 3 | 3 | N/A | VADJP_FB | see reg 592 | Bits 15:12 unused, Bits 11:0 DAC value
608 | 0260 | 0 | 0 | N/A | Event number LSBs | Bits 15:0 - low 16 bits of event number | 32-bit unsigned (with below)
609 | 0261 | 0 | 1 | N/A | Event number MSBs | Bits 15:0 - high 16 bits of event number | 32-bit unsigned (with above)
610 | 0262 | N/A | N/A | N/A | Seconds SST PLL locked | Number of seconds the PLL generating SST_IN has been locked | 16-bit unsigned

## IRS3B_Reg_Default
Register Map for IRS3B |  |  |  | 2013-03-27 | [6 empty] | 2500 | mV
[11 empty] | 4096 | counts
|  | Decimal equivalent at right --> | [8 empty] | 0.61035 | mV/count
[9 empty] | Default | Decimal | Voltage
"PCLK #" | Register/Value | Default value | # bits | [5 empty] | value (hex) | equiv. | equiv.
1 | Threshold Ch. 1 | TBD (0x000) | 12 | scan required |  |  |  |  | 0 | 0 | 0
2 | Threshold Ch. 2 | TBD (0x000) | 12 | [5 empty] | 0 | 0 | 0
3 | Threshold Ch. 3 | TBD (0x000) | 12 | [5 empty] | 0 | 0 | 0
4 | Threshold Ch. 4 | TBD (0x000) | 12 | [5 empty] | 0 | 0 | 0
5 | Threshold Ch. 5 | TBD (0x000) | 12 | [5 empty] | 0 | 0 | 0
6 | Threshold Ch. 6 | TBD (0x000) | 12 | [5 empty] | 0 | 0 | 0
7 | Threshold Ch. 7 | TBD (0x000) | 12 | [5 empty] | 0 | 0 | 0
8 | Threshold Ch. 8 | TBD (0x000) | 12 | [5 empty] | 0 | 0 | 0
9 | VBDbias | 0x400 | 12 | DAC buff bias for Vbias, Vbias2 |  |  |  | 0x400 | 400 | 1024 | 625
10 | Vbias | 0x380 | 12 | Buff amp bias -- initial sample |  |  |  | 0x380 | 380 | 896 | 546.875
11 | Vbias2 | 0x370 | 12 | Buff amp bias -- transfer samples |  |  |  | 0x370 | 370 | 880 | 537.1094
12 | Misc Reg (incl. SGN) | 0x000 | 8 |  |  |  |  | 0x000 | 0 | 0 | 0
13 | WBDbias | 0x400 | 12 | DAC buff bias for Wbias |  |  |  | 0x400 | 400 | 1024 | 625
14 | Wbias | 0x3D7 | 12 | Trigger Width adjust (Wbias) |  |  |  | 0x370 | 370 | 880 | 537.1094
15 | TCBbias | 0x400 | 12 | DAC buff bias for Trigger Comparator |  |  |  | 0x400 | 400 | 1024 | 625
16 | TRGbias | 0x350 | 12 | Trigger Comparator bias |  |  |  | 0x350 | 350 | 848 | 517.5781
17 | THDbias | 0x400 | 12 | DAC buff bias for Trigger Thresholds |  |  |  | 0x400 | 400 | 1024 | 625
18 | Tbbias | 0x400 | 12 | Internal Trigger Threshold buffer bias |  | (disable for external drive) |  | 0x400 | 400 | 1024 | 625
19 | TRGDbias | 0x400 | 12 | DAC buff bias for TRGbias2, TRGthref |  |  |  | 0x400 | 400 | 1024 | 625
20 | TRGbias2 | 0x350 | 12 | TRGbias for reference channel |  |  |  | 0x350 | 350 | 848 | 517.5781
21 | TRGthref | 0x800 | 12 | Trigger threshold for reference channel |  |  |  | 0x800 | 800 | 2048 | 1250
22 | Leading Edge SSPin | 0x060 | 8 | Timing Gen: SSPin |  |  |  | 0x060 | 60 | 96
23 | Trailing Edge SSPin | 0x010 | 8 |  |  |  |  | 0x010 | 10 | 16
24 | Leading Edge S1 | 0x028 | 8 | Timing Gen: S1 |  |  |  | 0x028 | 28 | 40
25 | Trailing Edge S1 | 0x058 | 8 |  |  |  |  | 0x058 | 58 | 88
26 | Leading Edge S2 | 0x068 | 8 | Timing Gen: S2 |  |  |  | 0x068 | 68 | 104
27 | Trailing Edge S2 | 0x018 | 8 |  |  |  |  | 0x018 | 18 | 24
28 | Leading Edge PHASE | 0x018 | 8 | Timing Gen: PHASE |  |  |  | 0x018 | 18 | 24
29 | Trailing Edge PHASE | 0x030 | 8 |  |  |  |  | 0x030 | 30 | 48
30 | Leading Edge WR_STRB | 0x040 | 8 | Timing Gen: WR_STRB | NOTE: WR_ADDR phase CRITICAL |  |  | 0x040 | 40 | 64
31 | Trailing Edge WR_STRB | 0x070 | 8 |  |  |  |  | 0x070 | 70 | 112
32 | Timing Generator Reg | 0x*** | 8 | Select Timing signals viewed, Phase clear, RCO running |  |  | decode for view state --See MONTIMING Page below
33 | PDDbias | 0x400 | 12 | DAC buff bias for CMPbias |  |  |  | 0x400
34 | CMPbias | 0x500 | 12 | Storage Cell Comparators (Pull-Down) bias |  |  |  | 0x500
35 | PUDbias | 0x400 | 12 | DAC buff bias for CMPbias |  |  |  | 0x400
36 | PUbias | 0xBF0 | 12 | Storage Column Comparators (Pull-Up) bias |  |  |  | 0xBF0
37 | SBDbias | 0x400 | 12 | DAC buff bias for Super Buffer bias |  |  |  | 0x400
38 | Sbbias | 0x400 | 12 | Super Buffer bias |  |  |  | 0x400
39 | ISDbias | 0x400 | 12 | DAC buff bias for ISEL |  |  |  | 0x400
40 | ISEL | 0x900 | 12 | Voltage Ramp Current |  | Initial config required |  | 0x900
41 | VDDbias | 0x400 | 12 | DAC buff bias for Vdly |  |  |  | 0x400
42 | Vdly | 0xB54 | 12 | Wilkinson counter adj voltage |  | Servo-lock target value |  | 0xB54
43 | VAPDbias | 0x600 | 12 | DAC buff bias for VadjP |  |  |  | 0x600
44 | VadjP | 0x5E8 | 12 | Timing Generator delay adjust PMOS |  | Initial config required |  | 0x5E8
45 | VANDbias | 0x600 | 12 | DAC buff bias for VadjN |  |  |  | 0x600
46 | VadjN | 0xAD0 | 12 | Timing Generator delay adjust NMOS |  | Servo-lock target value |  | 0xAD0
47 |  | undef.
48 |  | undef.
49 |  | undef.
50 |  | undef.
51 |  | undef.
52 |  | undef.
53 |  | undef.
54 |  | undef.
55 |  | undef.
56 |  | undef.
57 |  | undef.
58 |  | undef.
59 |  | undef.
60 | DatOut_ctr_CLR | AddrMode | ADDR_M | Clear Data Output Address Registers
61 | RD_ctr_CLR | AddrMode | ADDR_M | Clear Read Address Registers
62 | Start_WilkMon | AddrMode | ADDR_M | Start Wilkinson Reference counter
63 | Boin_CLR | AddrMode | ADDR_M | Stop/Clear Wilkinson Reference counter
64 | Trig_In | AddrMode | ADDR_M | Pulse Trigger test circuit

## IRS3B_TIMING_REGISTER
PCLK = 32 address decoding
Bit # | new name | when 0 | when 1 | signal name inside ASIC
0 | montiming_select(0) | see table at right | see table at right | nSel_0
1 | montiming_select(1) |  |  | nSel_1
2 | montiming_select(2) |  |  | nSel_2
3 | disable_RCOSSP | RCOSSP enabled inside ASIC | breaks link in ring oscillator | nRipSSP
4 | disable_RCOSST | RCOSST enabled inside ASIC | breaks link in ring oscillator | nRipSST
5 | select_SSP | SST or RCOSST brought out | SSP or RCOSSP brought out | nSSPSST
6 | select_SSX | RCO brought out to pin | SSX brought out | nRCO_SSXmon
7 | clear_phase | for normal running | clears flip-flop for PHAB signal | nCLR_PHASE
| montiming_select(2) | montiming_select(1) | montiming_select0) | montiming output
| 0 | 0 | 0 | A1
| 0 | 0 | 1 | B1
| 0 | 1 | 0 | A2
| 0 | 1 | 1 | B2
| 1 | 0 | 0 | PHASE
| 1 | 0 | 1 | PHAB
| 1 | 1 | 0 | SSPin
| 1 | 1 | 1 | WR_STRB

## I2C addresses
row | column | ASIC temperature | carrier eeprom | carrier GPIO2 | carrier GPIO1 | carrier GPIO0 | carrier DAC
|  | M18,N18 |  |  | AD4,AE1 |  | H13,F14 | I2C SCL,SDA nets
0 | 0 | 1001000 | 1010000 | 0111110 | 0111101 | 0111100 | 0010000
0 | 1 | 1001001
0 | 2 | 0111000 | [6 empty] | GPIOs are PCA9534A
0 | 3 | 0111001
1 | 0 | 1001010 | 1010001 | 0111010 | 0111001 | 0111000 | 0010001
1 | 1 | 1001011
1 | 2 | 0111010
1 | 3 | 0111011
|  | L10,K10 |  |  | AB4,P19 |  | E14,K14 | I2C SCL,SDA nets
2 | 0 | 1001000 | 1010000 | 0111110 | 0111101 | 0111100 | 0010000
2 | 1 | 1001001
2 | 2 | 0111000
2 | 3 | 0111001
3 | 0 | 1001010 | 1010001 | 0111010 | 0111001 | 0111000 | 0010001
3 | 1 | 1001011
3 | 2 | 0111010
3 | 3 | 0111011
|  |  |  | 0100000 | 0100001 |  |  | interconnect | GPIOs are PCA9534 | W18,J2
updates from 2013-03-01:
swapped GPIO0 and GPIO2 in this table from previous version to make it correspond to the schematic
the boxes with a red background are a mistake; solution is to exchange for PCA9534, then the address is 0100010
alternative is to treat GPIO2 on c13 as write-only device and ignore the conflict with the temperature sensor (which it will not be possible to read)

## DAC Mappings (IRS2 DC RevB2)
DAC (silkscreen label) | Channel | DAC Pin | Net Name | Old Default Value (hex) | Old Default Value (dec)
1 | 0 | 2 | TRG_THRESH01 | 76C | 1900
1 | 1 | 3 | TRG_THRESH23 | 76C | 1900
1 | 2 | 4 | VADJP | C9E | 3230
1 | 3 | 5 | VADJN | 42E | 1070
1 | 4 | 12 | TRGBIAS | 3E8 | 1000
1 | 5 | 13 | VBIAS | 44C | 1100
1 | 6 | 14 | TRG_THRESH45 | 76C | 1900
1 | 7 | 15 | TRG_THRESH67 | 76C | 1900
2 | 0 | 2 | TRGTHREF | 000 | 0
2 | 1 | 3 | ISEL | 7D0 | 2000
2 | 2 | 4 | SBBIAS | 640 | 1600
2 | 3 | 5 | PUBIAS | CE4 | 3300
2 | 4 | 12 | VDLY | AF0 | 2800
2 | 5 | 13 | CMPBIAS | 384 | 900
2 | 6 | 14 | PAD_G | 7FF | 2047
2 | 7 | 15 | WBIAS | 51E | 1310